## Slide 1: Quantum Sensors for Rare Event Searches
Matt Pyle
9/14/21
QIS Collaboration Meeting 2021

## Slide 2: 10 Years Ago: A Focus on WIMPs
Relic DM density suggests weak-scale cross sections
New physics (and particles) at the weak scale could solve the hierarchy problem
10-22 eV
keV
1019 GeV
10 GeV
WIMPS
Notes: Main idea: 10 years ago the majority of the experimental community was searching for WIMPs
DM Constraints:
10^-22 eV: size of the galaxy
< 100 eV: Fermions are degenerate
< 3 keV: warm DM limit
WIMPs: 10GeV < M < 100TeV (above 100TeV WIMP won’t freeze out)
Hiearchy problem:
- Why do different forces have such enormously different strengths
- Why is the Higgs mass so much lighter than the plank scale

## Slide 3: Today: Search for Dark Matter Everywhere
US Cosmic Visions: New Ideas in Dark Matter: 1707.04591
Many Well Motivated DM Models at Light Mass
10-22 eV
keV
1019 GeV
10 GeV
WIMPS
Dark Sectors: DM + New Mediators
100MeV
1meV
This Talk

## Slide 4: Kinematics: 2 Body Elastic Nuclear Scattering
Recoil Energy Scales as MDM2. Transfer of DM kinetic energy is really inefficient for elastic 2 Body Scatters when Mn >> MDM
DM
When Mn>> MDM

## Slide 5: Light Mass Dark Matter: Elastic Nuclear Scattering
<v>
𝛳=0
Energy Sensitivity is Primary Design Driver
1 meV
Yikes!

## Slide 6: Ionization Production in eV Scale Nuclear Recoils
<v>, 𝛳=0
For DM <~ 200MeV, no ionization expected from NR
Electronic Bandgap in Silicon
Frenkel Defect Dislocation Energy in Si: ~ 15 eV

## Slide 7: Coherent Vibrational Excitation Regime
Below O(10 MeV), we need to start thinking about DM nucleus interactions in terms of coherent vibrational mode production
Coherent Vibrational modes

## Slide 8: Phonon Sensitivity Curves
Zurek, Griffin, et al: 1910.10716
Exposure: 1kgyr
No Backgrounds
1910.10716
Heavy Mediator Single Phonon Sensitivity
Detectors with ~20meV single phonon sensitivity needed to really probe this space

## Slide 9: Single phonon sensitivity -> quantum sensing technology

## Slide 10: Everything Else Easier: Detector Size
Interaction Rate scales with 1/MDM
Light mass dark matter searches are tabletop!
LZ needs 10 tons to get to 10-47 cm2 at 100GeV, Light Mass DM searches only needs 1kg to reach the same level at 10MeV

## Slide 11
8B
PP
7Be
Smaller DM masses have less overlap with flat backgrounds!
PEP
Everything Else Easier: Flat Radiogenic Backgrounds
σ=10-43cm2

## Slide 12: Ultralight DM: Minimal Radiogenic Background Overlap with Signal
10-22 eV
keV
1019 GeV
10 GeV
WIMPS
Dark Sectors: DM + New Mediators
1meV
Experiments: DMRADIO, HAYSTAC, ADMX, ABRACADABRA, CASPER, … very vibrant scientific area
Heavy use of quantum sensing technology
But … minimal overlap with radiogenic backgrounds -> no planning for underground running
Unless your quantum sensor itself needs a low background environment

## Slide 13: SURF+ Quantum Sensing Sweet Spot: 1meV <MDM< 1GeV
10-22 eV
keV
1019 GeV
10 GeV
WIMPS
Dark Sectors: DM + New Mediators
100MeV
10meV
This Talk

## Slide 14: But … there seem to be lots of “Detector Backgrounds” <100eV that aren’t flat

## Slide 15: Quantum Sensors for Light Mass Dark Matter Work Program
1) Understand and Mitigate Background Rates in the 1meV-100eV range
2) Develop Ultra-Sensitive Vibration Technologies
2a) TES based
(SuperCDMS, SPICE, HeRALD)
2b) MKID+ Quantum Sensors

## Slide 16: 1a) Stress Related Microfracture Events

## Slide 17: Cryogenic Photon Detector (CPD) @SLAC
Sam
11g 4.5cm3 Si detector
TES based athermal phonon sensors (simply more mature)
sensitivity = 4eV (rms)

## Slide 18: Event Rate: Same @CUTE and SLAC
1-5keV: 1000 dru
x100 larger than CUTE background
x10 larger than any possible Radon contamination
Sam

## Slide 19: Rate Variation with time since cooldown
EDELWEISS (but similar qualitative behavior seen by CPD and CRESST)
From J. Billard

## Slide 20: Low Energy Excess Facts:
Same above and below ground -> not sourced by neutrons or gammas
Time since cooldown variation -> not sourced by neutrons or gammas
E > O(10 eV) -> can’t be long lived electronic states
Way worse for CPD than CDMS2 Si in 1-5keV range -> CPD exasperates the problem
Zero Yield:
  EDELWEISS (apples to apples)
  CDMSlite and CPD (not apples to apples)
Only hypothesis that fits these facts: stress induced micro-fractures

## Slide 21: Solution: Minimal Stress Holding Strategies
The smallest possible normal force is just the weight …. don’t clamp the crystal.
glued
resting

## Slide 22: Initial Results: Extremely Promising
We’ve shown that thermal stress in GE varnish produce 5Hz/cm2 of low energy microfractures
  Big Caveat: this is for GE varnish, not cirlex clamps!
Glued
Resting
Bill

## Slide 23: Complications: Increased Vibration Susceptibility
Resting support design significantly more susceptible to vibrational rubbing than cirlex clamps or glue 
  Seen at UCB and CUTE
  PT off data good
1cm2 Hanging Design
Siddhant/Roger
Bill
Option 1: Internal double mass vibration
decoupler (Sam)
Option 2: hanging design
  Idea: Nader & Bernard
  Google Sycamore Quantum Computer hangs from wire bonds
  Least amount of infrastructure change at CUTE

## Slide 24: 1b) EMI and IR Noise Mitigation

## Slide 25: Low Tc W Film Fabrication (TAMU)
What’s set W Tc? 2 crystal configurations
  Alpha: Tc=15mK
  Beta: Tc~ 3K
Goal: produce a stress free, alpha phase W film
Bouziane et al: Xe plasma produces better alpha films
  P. Brink found and developed work program
  CRESST did too!
TAMU:
  Upgraded SEGI and AJA with Xenon gas
  50 films made and Tc measured at LBL
  Tc=28mK film produced
Future work program
  Optimize for Tc, alpha, stress at the same time … we really want slightly compressed W (Xe) films @300K
  Reproducible 15<Tc<20mK
Appl. Phys. A 81, 209–215 (2005)

## Slide 26
25umx100umx40nm W TES at 59mK
Achieved Required Performance
  Bias Power (P0): 20fW for 25x100
  Falltime: 70us
  Estimated Resolution: 29meV (world leading)
PT off
x4 power ratio suggests ~<2fW parasitic power on lines … much smaller bias powers possible now!
23us falltime
TES Characterization and EMI Mitigation

## Slide 27: EMI Mitigation Improvements
Measure EMI shielding effectiveness with 200Hz modulation of 800MHz EMI … still have leakage
300K:
  Replace o-rings with Al gaskets -> this run
  Better filtering on fridge thermometry
200Hz amplitude modulation of 800MHz EMI
Make the 4K thermal shields into Faraday cages with line filtering
Steal copper powder filter from Martinis, Devoret, Clarke (PRB 35.4682 1987)

## Slide 28: Near Term Future: decrease volume/Tc
Sept ‘21: 28mK 25umx25umx40nm W TES (TAMU):
  Estimated P0 ~ 100aW -> x180 smaller than current performance
  Estimated ~ 1.5meV -> x20 smaller than current performance
Of course, it may take us a while to achieve this performance … but we’ll have the device to do it.

## Slide 29: Summary:
10meV-100MeV Dark Matter
  Theoretically well motivated
  Low’ish radiogenic backgrounds -> perfect for SURF
  Single phonon/photon sensitivity needed -> quantum sensors
  Small 100g-1kg
Work Program:
  Detector Backgrounds / Dark Counts
  Improve Sensing Technology

## Slide 30: Backup

## Slide 31: Optical Phonons
Acoustic:
Optical:
Due to their gapped nature, optical phonons are kinematically matched to IR and light mass DM!
Knapen, Lin, Pyle, & Zurek: 1712.06598
Momentum
Energy [meV]
Sapphire
1712.06598

## Slide 32: Dark Photon Couplings: Polar Crystals
Scattering via Light Dark Photon
Exposure: 1kgyr
1910.10716
In ionic crystals, optical phonons are oscillating electric dipoles!
Very large coupling to photons (black in the IR)… Very large coupling to the dark photons
30-100meV phonon sensitivities required

## Slide 33: Sapphire: Lot’s of Optical Phonon Bands
To search for thermal DM down to keV masses via scattering and ultracold bosonic DM to 30 meV, we need a detector sensitive to a single optical phonon
Sapphire is a complex crystal with 10 atoms in its unit cell. Dark Matter can interact with lots of different modes
Dark Photon Dark Matter Absorption
1807.10291
1807.10291

## Slide 34: Other Signal Channels for Interacting with Light Mass Dark Matter

## Slide 35: Inelastic e- Recoils in Semiconductors
Energy [eV]
Momentum
Band Diagram for Si
Detector Requirement: Sensitivity to single e/h pairs (1 eV ) with negligible dark count rate
<v>
vesc
e- excitation momentum and energy scales in semiconductors well matched to 1 MeV-100MeV DM
Essig et al: 1108.5383

## Slide 36: Inelastic e- Recoils: Scintillating Crystals
Use a low bandgap scintillating crystal (GaAs, NaI) and couple to a single photon sensitive large area detector with no dark count rate
Penalty: Scintillation Production Efficiency
Derenzo et al: 1607.01009

## Slide 37: Optical/IR Haloscope
Baryakhtar, Huang, Lasenby
1803.11455
Momentum matching via multilayer stack
Requirements: Single photon sensitivity with no dark counts

## Slide 38: Building Detectors with Sensitivity to Single Phonons, Single IR Photons, and Single e/h (without Luke-Neganov Gain)

## Slide 39: 100meV Excitation Sensitivity -> Low Temperature Detectors
300K = 26meV
To sense 100meV excitations we’ll need to cool the detector down to near absolute zero
Use superconducting detector technology
  MKIDs
  Transition edge sensors
  SNSPD (only for photon applications)

## Slide 40: The Simplest Thermal Calorimeter
R
T
Transition Edge Sensor (TES):
  A superconducting metal film (W) that is externally biased so as to be within its superconducting/normal transition
  Must use low Tc and very small volume TES -> hard to get gram-day exposures when your TES (25umx25umx40nm) is 500fg … only directly useful for IR Haloscope

## Slide 41
R&D: ultra sensitive TES
Build and test simple TES test structures for noise is performance
Tests at SLAC SuperCDMS test facility
2004.10257
Tungsten TES
  Tc= 41mK
  40nm thick

## Slide 42: 100um x400um TES Characterization
TES Power vs TES Bias Voltage
TES Response to Square Wave Jitter
Normal Resistance: 630mOhm
Bias Power: 35fW
  Your average CMB TES: 1 pW
  Slight calibration error
Complex Impedance
  Simple 2 pole TES dynamical model perfectly fits response
TES falltime: ~66us (2.4kHz)
  Relatively fast
Bias Power (pW)
Bias Voltage (uV)
Time (us)
TES Current (uA)
2004.10257
2004.10257

## Slide 43: 100um x400um TES Noise
Noise matches theory to within x2.5
  Sptot = 1.5x10-18 W/rtHz
Relatively high sensor bandwidth (2.4kHz)
𝛔E = 40 meV
Lots of environmental noise mitigation work to do
2004.10257

## Slide 44: More Sensitivity -> Decrease Volume
Volume (40nm thick) | Bias Power | Expected Sensitivity
25x25um2 | 0.5 fW | 5meV
100x25um2 | 2.2 fW | 10meV
200x50um2 | 9 fW | 20meV
400x100um2 | 35 fW | 40meV
200x500: barely operates
100x25um: completely normal
We have 5 fW of DC environmental parasitic power hitting our TES. Our current primary challenge is to continue to improve environmental isolation!
2004.10257

## Slide 45: Faraday Cage #1 -> Dilution Fridge
If your E&M signal can get out, the environmental EMI can get in!
Need to carefully filter all signal lines breaching the faraday cage
3MHz Pi Filters
ISO K Al gaskets
(Not rubber o-ring)
Eccosorb reflectivity
Eccosorb AN79

## Slide 46: Faraday Cages 2 & 3: 4K & 10mK
Inner thermal shields would act as additional Farday Cages … if there were filters on all the lines.
Bock Black for IR light leaks
Steal copper powder filter from Martinis, Devoret, Clarke (PRB 35.4682 1987)
SENSORS
Bock Black
Bock Black
Bock Black

## Slide 47: Phonon Detectors

## Slide 48: 2nd most simple thermal calorimeter
Couple the sensor to a large volume insulator -> low heat capacity

## Slide 49: Problem: Decoupling between the Sensor and Absorber
As T is decreased, it’s harder and harder to keep the sensor thermally coupled to the absorber
Energy leaks out of the absorber through Gab before its measured
TES sensitive to power fluctuations through Gtb
Physical clamps that support the absorber
Kapitza boundary conductance scale as as T3
e-/phonon thermal conductance scales as T4

## Slide 50: Athermal Phonon Sensor Technology
Collect and Concentrate Athermal Phonon Energy into Sensor
R
T

## Slide 51: Excitation Detectors & Volume Scaling
PMT
Scintillator
Reflector
Will these detectors have the same energy sensitivity?
Yes, if:
Lifetime of the athermal excitation (photon) is really long
Excitation absorption dominated by sensor
Position Sensitivity
PMT
Scintillator

## Slide 52: Optimizing the Athermal Phonon Excitation Detectors
TES
Absorber
TES
Absorber
Minimize the number/volume of the TES sensors instrumented on the surface to the point that you begin to see the bare surface thermalization rate

## Slide 53: Athermal Phonon Sensor Sensitivity Scaling
Power Noise for various G
Phonon Power Noise [W/rtHz]
G
Lower 𝑣sensor (lower Tc) if 𝑣signal < 𝑣sensor
Lower 𝑣signal (decrease Al coverage) if 𝑣signal > 𝑣sensor
You can always say on VTc3 scaling (in principle)
45mK-> 10mK: 2eV -> 20meV

## Slide 54: Athermal Phonon Thermalization at Surfaces
Athermal phonon surface thermalization probability found to depend upon
  Crystal
  Surface roughness
  Surface cleanliness
(W. Knaak et al, Phonon Scattering in Condensed Matter V,1986)
0.1%-1% of the crystal surface covered with athermal phonon sensors … 1/1000-1/100 thermalization probability needed
Si, Ge -> ok

## Slide 55: Step 2: Make the Athermal Phonon Sensor
Measured Phonon Collection Efficiency = 20% -> 40% (theoretical limit)
R&D Work Plan
  Optimize Collector/TES (W/Al) interface
  Improve quasi-particle trapping in collector fin

## Slide 56: Step 3: Fabricate Sensors on Crystal

## Slide 57: Athermal Phonon R&D
Matt Pyle
For
T. Aramaki, P. Brink, J. Camilleri, C. Fink, R. Harris, Y. Kolomensky, R. Mahapatra, N. Mirabolfathi, R.Partridge, M. Platt, B. Sadoulet, B. Serfass, S. Watkins, T. Yu
2007.14289

## Slide 58: Cryogenic Photon Detector (CPD)
  3” diameter Si wafer (45.6 cm^2)
  1mm thick
  Distributed athermal phonon sensors minimize phonon collection time (as fast as it can be for its size)
    Athermal Phonon collection time estimated to be ~20us
    2.5% sensor coverage
2007.14289

## Slide 59: Measured Performance: Tc & IV
  Tc= 41.5mK
  IV curves show that the detector and electronics are behaving well
    Rn = 88 mOhms (300mOhm Expected … TES too wide!)
    Rp = 8 mOhms
    Bias Power (P0) = 3.9 pW (1.4pW expected)
    Rn does not vary as a function of QET bias
Bias power flat in transition
2007.14289

## Slide 60: Measured Performance: dIdV
TES sensor pretty fast @ 60us. However, it’s not as fast as the estimated athermal phonon collection (20us)
Therefore, we expect phonon signals to have a 20us rise time (athermal phonon collection) and a 60us fall time. Seen for low energy comptons in average pulse shape!
Pulse shape varies with energy due to local TES saturation.
2007.14289

## Slide 61: Measured/Theoretical Noise
Noise ~x1.5 worse than theoretical expectations
Lots of lines from pulse tube and turbo pump
TFN seems too large …
2007.14289

## Slide 62: Integral Estimators for relative 55Fe calibration
Since pulse shape has significant variation with energy, we must use noisy but minimally biased DC estimators to fit the 55Fe calibration lines
Al Fluoresence @ 1.49 keV
55Fe K𝛼
@5.9keV
17% Athermal Phonon Collection Efficiency
2007.14289

## Slide 63: Calibrating Pulse Shape Dependent Energy Estimators to the DC estimator (Pulse Tube On)
𝜎E =3.9 ± 0.25 eV (stat)-0.7 (sys)
Baseline Resolution
+0.0
2007.14289

## Slide 64: First Dark Matter Search
2007.14289
In collaboration with SuperCDMS, we ran the CPD detector at the SLAC surface test facility
  Significantly limited by cosmogenic backgrounds
  10gd exposure
World leading DM sensitivity from 87-140 MeV
Just the beginning …

## Slide 65: First Dark Matter Search
20 gram-day DM search completed at SLAC in collaboration with SuperCDMS … results any day
Now Running in SNOLAB CUTE Facility
2007.14289

## Slide 66: Prototype Design Estimated Sensitivities
# TES | 100
TES Dimensions | 50um x 2um x40 nm
TES Rn | 320mOhm
Fin Length | 125um
W/Al Overlap | 15um
Fractional Al Coverage | 1%
Tc | 40mK
Bias Power | 48fW
Power Noise | 5.1e-19 W/rtHz
Phonon absorption time | 106us
Sensor fall time | 97us
Collection efficiency | 19%
σE | 219 meV
New 1cm3 Prototype Test Design
With a Si Absorber: single e/h sensitivity without Luke-Neganov gain. Can be used for inelastic electronic recoil DM
World Leading Elastic Nuclear Recoil DM search potential

## Slide 67: Summary:
Far IR photon detectors, and Athermal Phonon Small Volume detectors are a promising technique to search for 10meV-100meV Dark Matter
Current Progress:
  3” large are photon detector 3.9eV resolution
  100x400um TES test chip: 40meV resolution
  Running photon detector at CUTE soon
R&D Plan: Towards single phonon/single phonon sensing
  Decrease environmental noise
  Increase TES sensitivity: lower Tc
  Fabricate on a variety of crystal substrates (SiO2)
  Study how phonon surface themalization depends on surface roughness.
  Improve Collection efficiency: optimize W/Al interface